## Main
DISCLAIMER: All prices are approximate and subject to change by the vendors without prior notice
Part # | Vendor | Description | QTY | Unit Price (USD) | Price (USD) | Assembly | Subassembly | Comment
Camera
01-IRIS-15-PCIe-M-16-C | Photometrics | IRIS-15-PCIe sCMOS camera | 1 | 12350 | 12350 | Detection arm |  | FOV 25 mm diagonal
Excitation laser combiner options
Recommended laser combiner specs: 3-4 lasers: (405 optional), 488, 561, 638 nm, power 50-150 mW. Digital and analog power modulation. Coupling to SM fiber with FC/APC end connector.
| Oxxius | L4Cc 488-561-638 nm, 100 mW each, 2 fiber ports, switching module MDL-FSTM | 1 | 25000 | 25000 | Excitation |  | Includes switching module between 2 laser ports, controlled by TTL signal. The 405nm is NOT included!
Sample & Microscope Stages
Dual-LS100-FTP | ASI | Two parallel LS-100 servo actuators, table mounting brackets not included | 0 | 6500 | 0 | Sample motion |  | Sample Y stage
LS-50-AMCCH | ASI | LS-50 Linear Stage, with sub-micron accuracy, and 50 mm of travel | 0 | 2250 | 0 | Sample motion |  | Sample X stage
LS-5012 | ASI | Plate for attaching flat surface of linear stage to breadboard. 3.9" x 5.9" | 0 | 160 | 0 | Detection arm | Detection path
mesoSPIM-Bridge-wide | Mechanical shop or ASI | Bridge part between two parallel LS-100 servo actuators | 1 | 160 | 160 | Sample motion |  | Bridge for XZR stage assembly. NOT included in the STAGE-4D-MESOSPIM-K1
LS-50-AMCLH | ASI | LS-50 Linear Stage, with sub-micron accuracy, and 50 mm of travel | 0 | 2250 | 0 | Detection arm | Detection path | Focus (F) stage, moves camera and detection lens assembly
DV-6010-C | ASI | Dovetail mount pair (VTS-2112 & CDZ-1010) 60 x 62 x 10 mm overall dimensions. | 0 | 230 | 0 | Sample motion |  | Cross-coupling X and Z stage adapter
LS-50S-AMECH | ASI | LS-50 Linear Stage, with sub-micron accuracy, 50 mm of travel, and stainless steel slide | 0 | 2425 | 0 | Sample motion |  | Sample Z-stage
LE-LS-HH-70 | ASI | Encoder Linear LE Heidenhain 70mm scale installed on AMECHS stage | 0 | 3000 | 0 | Sample motion |  | Sample Z-stage encoder
C60-3060-CMR-1011 | ASI | LS-50 slide to C-mount adapter | 0 | 150 | 0 | Sample motion |  | Rotary stage adapter
C60-3060-SRS | ASI | Rotary stage with collet for holding a sample. | 0 | 3000 | 0 | Sample control |  | Sample Rotation (R) stage
STAGE-4D-MESOSPIM-K1 | ASI | All stage parts above offered by ASI | 1 | 19965 | 19965 |  |  | Bundle of 6 stages with mounting hardware
TG8-BASIC | ASI | Controller: Includes TG8, TGCOM, 2x TGDCM2, SA-JOY+ZF Contols four DC servo motors, | 0 | 5800 | 0 | Motion control |  | Motion controller with pre-installed cards
TGDCM2 | ASI | Dual Axis motion control card - one slot | 0 | 1100 | 0 | Motion control |  | Motion controller card
TG8-MESOSPIM-K1 | ASI | The controller bundle offered by ASI | 1 | 6900 | 6900 |  |  | Bundle of motion controller
AB90H | Thorlabs | Slim Right-Angle Bracket with Counterbored Slots | 4 | 27 | 108 | Sample motion |  | Table mounting for the Y axis (two parallel LS-100 servo actuators Dual-LS100-FTP)
Sample handling
KB25/M | Thorlabs | Complete 25 mm x 25 mm Kinematic Base, Top and Bottom Plates, M4 Counterbores | 2 | 70 | 140 | Sample handling |  | Magnetic bases to attach sample holders and immersion cuvette
KBB25/M | Thorlabs | Bottom Plate Only of the KB25/M Kinematic Base, M4 Counterbore | 6 | 36 | 216 | Sample handling |  | Sample holder side (nests, not balls) of the kinematic base
RLA150/M | Thorlabs | Dovetail Optical Rail, 150 mm, Metric | 1 | 43 | 43 | Sample handling |  | Immersion cuvette base
RC2/M | Thorlabs | Dovetail Rail Carrier, 50.8 mm x 25.4 mm, M6 Counterbore, M4 Taps | 1 | 29 | 29 | Sample handling |  | Rail carrier for immersion cuvette base
RD-PDT-S | Radiant Dyes | Prism Turn Table Small | 1 | 195 | 195 | Sample handling |  | Base for immersion cuvette
RD-PDT-S-to-RC2-height40mm.ipt | 3D printed | Adapter for mating Radyan Dyes prism turn table to Thorlabs rail carrier RC2 | 1 | 10 | 10 |  |  | See file RD-PDT-S-to-RC2-height40mm.ipt (.stl) for details.
Computer
Recommendations: at least 64 GB RAM, 2 TB SSD drive, 2 PCIex8 slots
|  | Imaging Workstation with Screen | 1 | varies | 0
Excitation arms
33-494 | Optotune | Optotune EL-16-40-TC-VIS-5D-1-C lenses | 2 | 780 | 1560 | Excitation | ETL Unit
88-940 | Optotune | Optotune Lens Driver EL-E-4i | 2 | 280 | 560 | Excitation | ETL Unit | Have to be modified: BNC cables need to be soldered: see https://github.com/mesoSPIM/mesoSPIM-hardware-documentation/wiki/mesoSPIM_preparing_ETL_drivers
88-941 | Optotune | Optotune 6-way Hirose Male/Female Cable, 1m Length | 2 | 85 | 170 | Excitation | ETL Unit
SM1L10 | Thorlabs | SM1 Lens Tube, 1.00" Thread Depth, One Retaining Ring Included | 2 | 13 | 26 | Excitation | ETL Unit
SM1M20 | Thorlabs | SM1 Lens Tube Without External Threads, 2" Long, Two Retaining Rings Included | 2 | 17 | 34 | Excitation | ETL Unit
SM1A10 | Thorlabs | Adapter with External C-Mount Threads and Internal SM1 Threads, 4.4 mm Spacer | 2 | 18 | 36 | Excitation | ETL Unit
varies | Digitec or elsewhere | Powered USB-Hub | 1 | 30 | 30 | Excitation | ETL Unit | Simple USB-Hub for powering the ETL-Drivers
91863 | Nikon-modified | Nikon AF-S 50 mm f/1.4 G Objective | 2 | 400 | 800 | Excitation | Excitation obj | Has to be modified
LCP06M-Nikon-modified.ipt | Mechanical workshop | Nikon objective mount using Thorlabs LCP06/M 60 mm Cage Plate with Ø2" Double-Bore Optic Mount | 2 | 300 | 600 | Excitation | Excitation obj mount | See /custom-parts/excitation-objective/LCP06M-Nikon-modified.ipt
LCP06/M | Thorlabs | 60 mm Cage Plate with Ø2" Double-Bore Optic Mount, M4 Tap | 0 | 40 | 0 | Excitation | Excitation obj mount | Goes into LCP06M-Nikon-modified. If DIY, lathe turning is required.
SM2NFM | Thorlabs | Adapter with External SM2 Threads and Nikon Female F-Mount Ring | 0 | 92 | 0 | Excitation | Excitation obj mount | Goes into LCP06M-Nikon-modified. If DIY, the F-mount steel ring must be taken out and mounted into the modified LCP06M plate using screws.
Cage60mm-2sides-InnerD75mm-L75mm.stl | 3D printed | Laser safety snap-on covers | 2 | 1 | 2 | Excitation | Excitation obj mount | See files files in folder /custom-parts/laser-safety/
SM1FCA | Thorlabs | FC/APC Fiber Adapter Plate with External SM1 (1.035"-40) Threads, Wide Key (2.2 mm) | 2 | 31 | 62 | Excitation | Laser input | One per arm
CP33/M | Thorlabs | SM1-Threaded 30 mm Cage Plate, 0.35" Thick, 2 Retaining Rings, M4 Tap | 10 | 16 | 160 | Excitation | Cage assembly | 5 per arm
ER8-P4 | Thorlabs | Cage Assembly Rod, 8" Long, Ø6 mm, 4 Pack | 2 | 41 | 82 | Excitation | Cage assembly | 8 rods per arm
AC254-060-A-ML | Thorlabs | f=60 mm, Ø1" Achromatic Doublet, SM1-Threaded Mount, ARC: 400-700 nm | 2 | 99 | 198 | Excitation | Laser collimator | 1 per arm, if laser combiner has 2 fiber outputs (Oxxius)
AC254-100-A-ML | Thorlabs | f=100 mm, Ø1" Achromatic Doublet, SM1-Threaded Mount, ARC: 400-700 nm | 4 | 104 | 416 | Excitation | A 1:1 relay from ETL to galvo | Lenses L2, L3 in the design view
KCB1E/M | Thorlabs | Right-Angle Kinematic Elliptical Mirror Mount with Tapped Cage Rod Holes, 30 mm Cage System and SM1 Compatible, M4 and M6 Mounting Holes | 4 | 189 | 756 | Excitation | Laser folding | M1, M2 fold mirrors, in each arm
BBE1-E02 | Thorlabs | 1" Broadband Dielectric Elliptical Mirror, 400 - 750 nm | 4 | 110 | 440 | Excitation | Laser folding | M1, M2 fold mirrors, in each arm
PH150/M | Thorlabs | Ø12.7 mm Post Holder, Spring-Loaded Hex-Locking Thumbscrew, L=150 mm | 4 | 13 | 52 | Excitation | Mechanical support | Optical cage support, 2 per arm
PH75/M | Thorlabs | Ø12.7 mm Post Holder, Spring-Loaded Hex-Locking Thumbscrew, L=75 mm | 2 | 8 | 16 | Excitation | Excitation obj mount
PH40/M | Thorlabs | Ø12.7 mm Post Holder, Spring-Loaded Hex-Locking Thumbscrew, L=40 mm | 4 | 7 | 28 | Excitation | Mechanical support | Optical cage support, 2 per arm
BA1S/M | Thorlabs | Mounting Base, 25 mm x 58 mm x 10 mm | 6 | 5 | 30 | Excitation | Mechanical support | Clamping post holders to optical table
TR100/M | Thorlabs | Ø12.7 mm Optical Post, SS, M4 Setscrew, M6 Tap, L = 100 mm | 4 | 6 | 24 | Excitation | Mechanical support | Optical cage support, 2 per arm
TR40/M | Thorlabs | Ø12.7 mm Optical Post, SS, M4 Setscrew, M6 Tap, L = 40 mm | 4 | 5 | 20 | Excitation | Mechanical support | Optical cage support, 2 per arm
TR75/M | Thorlabs | Ø12.7 mm Optical Post, SS, M4 Setscrew, M6 Tap, L = 75 mm | 2 | 5 | 10 | Excitation | Excitation obj mount
XRN25P | Thorlabs | Compact 25 mm Travel Linear Translation Stage, Side Micrometer, 1/4"-20 Taps | 2 | 398 | 796 | Excitation | M3 mirror assembly | Linear stage for M3 mirror adjustment
mirror-bracket.ipt | Mechanical workshop | Braket for M3 kinematic mirror mount | 2 | 150 | 300 | Excitation | M3 mirror assembly | Bracket for M3 mirror mounting on linear stage. Folder /custom-parts/excitation-mirror-bracket/
XRN-B1/M | Thorlabs | Base Plate for Stages with 2" Dovetails, Metric Slot Spacing | 2 | 30 | 60 | Excitation | M3 mirror assembly | Adapter plate for linear stage
QS15-mount-modified-Thorlabs-CagePlate-LCP03.ipt | Mechanical workshop | Modified 60 mm Blank Cage Plate | 2 | 300 | 600 | Excitation | Galvo assembly | See /custom-parts/galvo-assembly/15mm-galvo(Thorlabs-QS15)-mount/QS15-mount-modified-Thorlabs-CagePlate-LCP03.ipt
LCP31/M | Thorlabs | 60 mm Blank Cage Plate, M4 Tap (supersedes obsolete LCP03) | 0 | 36 | 0 | Excitation | Galvo assembly | Included in ASI bundle. If DIY, needs modification (machining), see /custom-parts/galvo-assembly/15mm-galvo(Thorlabs-QS15)-mount/QS15-mount-modified-Thorlabs-CagePlate-LCP03.ipt
QS15X-AG | Thorlabs | Ø15 mm Beam Galvo System, X-Axis Protected Silver Mirror | 2 | 2104 | 4208 | Excitation | Galvo assembly | Standard cables are 60 cm long.
CBLS3F | Thorlabs | Command and Power Cables for QS15/20/30/45, SS, & SP Series Galvo Scanners | 2 | 82 | 164 | Excitation | Galvo assembly
GPS011-EC/US/JP | Thorlabs | 1D or 2D Galvo System Linear Power Supply, 110, 230, or 100 VAC (EC, US, JP option) | 1 | 520 | 520 | Excitation | Galvo assembly | Preferred power supply for 2 galvos. Needs a simple adapter (3 wires), see Wiki.
GPWR15 (unsafe) | Thorlabs | ±15 V Power Supply for BLINK & QS, SS, SP, & XG Series Galvo Scanners | 0 | 390 | 0 | Excitation | Galvo assembly | Non-preferred (but offered by Thorlabs) Power supply for 2 galvos. Bare wires sticking around, unsafe!
HeatSink-sheet-H2mm.pdf | Mechanical workshop | Galvo heat sink assembled from sheets, each galvo requires 7 sheets | 14 | 20 | 280 | Excitation | Galvo assembly | Claser-cut or water-jet cut from 2-mm Aluminum. See folder /custom-parts/galvo-assembly/15mm-galvo(Thorlabs-QS15)-mount/
TODO | Mouser | Heat sink for galvo driver | 1 | 10 | 10 | Excitation | Galvo assembly | TODO: find specific model
Detection: Camera and lens mount
LS-50-AMCLH | ASI | LS-50 Linear Stage, with sub-micron accuracy, and 50 mm of travel | 1 |  |  |  |  | see Sample and Microscopy Stages
XT66SD-250 | Thorlabs | 6 mm Single Dovetail Rail with Mounting Counterbores, L = 250 mm | 1 | 63 | 63
XT66P2/M | Thorlabs | Rail Carriage for 66 mm Rails with M4 & M6 Taps | 2 | 69 | 138
LS-5012 | ASI | Plate for attaching flat surface of linear stage to breadboard. 3.9" x 5.9" | 1 |  |  |  |  | see Sample and Microscopy Stages
Detection: Filter wheel
5-slot filter wheel for 25-mm filters
ZWO EFW-MINI | ZWO | ZWO EFW Mini 5 x 31 mm | 1 | 200 | 200 |  |  | Check local distributors (astronomy shops)
25mmFilterAdapter.stl | 3D printed (SLS) | Adapter to hold 25 mm filters in 31 mm slots | 5 | 1 | 5 |  |  | See /custom-parts/detection-accessories for printable 3D model
SM2L05 | Thorlabs | 0.5" long SM2 lens tube | 1 | 26 | 26 |  |  | Adapter: Objective to filter wheel, via M2M40S thread adapter
SM2T2 | Thorlabs | SM2 (2.035"-40) Coupler, External Threads, 1/2" Long | 1 | 37 | 37 |  |  | SM2 coupler to rotate the filter wheel and fix it at desired angle the with retaining ring
TMA4 | Thorlabs | Adapter with External T-Mount to Internal SM2 | 2 | 24 | 48 |  |  | Adapters for both sides of the ZWO EFW Mini filter, which has T-mount threads (M42)
Filters
Filters have to be selected according to lasers & needs (flurescent markers and dyes used)
F57-405 | AHF | QuadLine Rejectionband ZET405/488/561/640 in 25 mm diameter | 1 | 625 | 625 | Detection path |  | A frequently used quad-band filter, allows all fluorescence through
Detection: Objectives
18-402 | Edmund Optics | Mitutoyo Plan Apo Brightfield and Darkfield, 2X/0.055 | 1 | 912 | 912 | Detection path |  | Works better with Olympus SWTLU-C tube lens
18-403 | Edmund Optics | Mitutoyo Plan Apo Brightfield and Darkfield, 5X/0.14 | 1 | 837 | 837 | Detection path |  | Works best with the Mitutoyo MT-1 tube lens
18-404 | Edmund Optics | Mitutoyo Plan Apo Brightfield and Darkfield 7.5X/0.21 | 0 | 1540 | 0 | Detection path |  | Works best with the Mitutoyo MT-1 tube lens. Optional.
18-405 | Edmund Optics | Mitutoyo Plan Apo Brightfield and Darkfield 10x/0.28 | 1 | 1026 | 1026 | Detection path |  | Works best with the Mitutoyo MT-1 tube lens
378-847 | Mitutoyo | Mitutoyo G Plan APO 20X/t3,5 | 0 | 3670 | 0 | Detection path |  | Corrected for 3.5 mm of glass (n1.52). Thread is unusual, but fits SM1.
Mitutoyo-BD-black-ring.ipt | 3D printed | Back cover for the Mitutoyo BD objectives | 5 | 1 | 5 | Detection path |  | Only needed for BD-series objectives.
M2M40S | Thorlabs | SM2 adapter for Mitutoyo objectives | 1 | 60 | 60 | Detection path
Detection: Tube lens
36-401 | Edmund Optics | Mitutoyo MT-1 Accessory Tube Lens | 1 | 752 | 752 | Detection path |  | Preferred choice for Mititoyo objectives, except the 2X/0.055 one
Adapter-MitutoyoMT1-SM2.stl | 3D printed | Adapter for SM2 lens tube, attached to SM2 retaining ring | 1 | 1 | 1 | Detection path
SM2E60 | Thorlabs | SM2 Series Extension Tube, 6" | 1 | 65 | 65 | Detection path |  | Tube lens is inserted inside this tube
SM2M20 | Thorlabs | SM2 Lens Tube Without External Threads, 2" Thread Depth | 1 | 32 | 32 | Detection path |  | Extends the SM2E60 and mounts to the SM2NFM2
SM2NFM2 | Thorlabs | Adapter with External SM2 Threads and Male F-Mount Ring | 1 | 112 | 112 | Detection path |  | Adapter to camera with F-mount
Electronics / Signal generation
778512-01 | National Instruments | NI PXI-6733 with 8 16-Bit Waveform Analog Outputs | 1 | 3582 | 3582 | Signal generation |  | The signal generation card, 8 analog and 8 digital channels. Potentially replaceable with two NI-PCIe-6323 cards (and no chassis), but this needs testing and possibly software changes.
781163-01 | National Instruments | NI PXIe-1073 Integrated MXIe, 5 Periph Slots | 1 | 1630 | 1630 | Signal generation |  | Chassis to mount the signal generation card
779504-01 | National Instruments | NI PCIe-8361, 1 Port MXI-Express Interface | 1 | 886 | 886 | Signal generation |  | Connects the Chassis to the PC. Can be replaced with more modern two-port PCIe-8362 for a higher price.
779500-03 | National Instruments | MXI-Express Cable, Gen 1 x1, Copper, 3m | 1 | 163 | 163 | Signal generation |  | Cable between the Chassis and the PC
763065-01 | National Instruments | Power Cord, 220V, 10A (plug country-specific) | 1 | 11 | 11 | Signal generation
777643-01 | National Instruments | BNC-2110 Noise Rejecting, Shielded BNC Connector Block | 1 | 520 | 520 | Signal generation
199006-01 | National Instruments | SH68-68-EPM Shielded Cable, 68 D-Type to 68 D-Type, 1M | 1 | 193 | 193 | Signal generation
PAA236R | Thorlabs | SMB Coaxial Cable, 90° SMB Female to BNC Male, 36" (914 mm) | 10 | 12 | 120 | Signal generation |  | Depending on the laser control interface (Oxxius: SMB).
CA2806 | Thorlabs | SMA Coaxial Cable, SMA Male to BNC Male, 6" (152 mm) | 1 | 14 | 14 | Signal generation |  | Depends on laser interface. Oxxius: IO connector must be modified, pin24 (MDL-FSTM) and pin16 (GND Digital), to accept arm switching signal.
2249-C-120 | Thorlabs | 0.3 m BNC Cable | 4 | 15.25 | 61 | Signal generation
2249-C-24 | Thorlabs | 0.6 m BNC Cable | 10 | 16.5 | 165 | Signal generation
2249-C-60 | Thorlabs | 1.5 m BNC cable | 10 | 24 | 240 | Signal generation
2249-C-120 | Thorlabs | 3m BNC cable | 6 | 24 | 144 | Signal generation
General lab supplies
LAB-16 | Newport | Ultra Jet Canned Air Duster, 10 oz. (300ml) | 4 | 16.5 | 66 |  |  | Or other canned duster
MC-50E | Thorlabs | Lens Tissues, 25 Sheets per Booklet, 50 Booklets in a Closeable Box | 1 | 85.1 | 85.1
BKF12 | Thorlabs | Black Aluminium Foil | 2 | 32 | 64 |  |  | Sometimes useful
T743-2.0 | Thorlabs | High-Performance Black Masking Tape, 2" x 180' (50 mm x 55 m) Roll | 2 | 51.58 | 103.16
PSS7 | Thorlabs | 7-Piece Precision Screwdriver Set | 1 | 60 | 60 |  |  | Or other, cheaper sets
HW-KIT6/M | Thorlabs | Imperial Screw Thread Adapter Kit | 1 | 136.48 | 136.48
HW-KIT6/M | Thorlabs | Metric Screw Thread Adapter Kit | 1 | 136.48 | 136.48
HW-KIT1/M | Thorlabs | M4 Cap Screw and Hardware Kit | 1 | 60.95 | 60.95
HW-KIT3/M | Thorlabs | M4 Setscrew and Hardware Kit | 1 | 119.416 | 119.416
HW-KIT4/M | Thorlabs | M6 Setscrew and Hardware Kit | 1 | 105.8 | 105.8
HW-KIT2/M | Thorlabs | M6 Cap Screw and Hardware Kit | 2 | 119.6 | 239.2
HW-KIT5 | Thorlabs | 4-40 Cap Screw and Hardware Kit | 1 | 100 | 100
TC3/M | Thorlabs | 15-Piece Balldriver and Hex Key Kit with Stand, Metric | 2 | 92 | 184
TC2 | Thorlabs | 20-Piece Balldriver and Hex Key Kit with Stand, Imperial | 1 | 96.23 | 96.23
SPW801 | Thorlabs | Adjustable Spanner Wrench | 1 | 102.44 | 102.44
SPW602 | Thorlabs | SM1 Spanner Wrench, Graduated, Length = 3.88" | 2 | 28 | 56
|  | Multiport-power outlets | 4
|  | Cable wraps | 1 |  |  |  |  | Lots of them
|  | Isopropanol p.a. for lens cleaning | 1
Sample cuvettes | Material: glass, RI=1.52. For quartz (e.g. CLARITY imaging, RI=1.47), replace UG with UQ in the part name.
| 3D printer or mechanical workshop | Large-cuvette-mount-40mm-V3.ipt | 4 | 40 | 160 | Immersion chamber mount | Machined from POM (Polyoxymethylen) or 3D printed from PA12 (polyamide, nylon) | One for each planned immersion cuvette mount / imaging medium
| Portmann Instruments | UG-753-H75 40x40x75 fire-fused | 2 | 331 | 662 | Immersion chamber
| Portmann Instruments | UG-203 5x10x45 | 4 | 25 | 100 | Immersion chamber
| Portmann Instruments | UG-204 10x10x45 | 4 | 28 | 112 | Immersion chamber
| Portmann Instruments | UG-205 10x20x45 | 6 | 121.5 | 729 | Immersion chamber
|  |  |  | Total: | 94317.256 | USD
Machined parts
Part name | Qty | Notes
HeatSink-sheet-H2mm (from 2mm Alu sheet) | 14 | made from 2mm Alu sheets: laser cutter, waterjet, or EDM
LCP06M-Nikon-modified | 2 | machined from Thorlabs LCP06M
mirror-bracket | 2 | machined from stock Alu
QS15-mount-modified-Thorlabs-CagePlate-LCP03 | 2 | machined from Thorlabs LCP03
3D printed parts from PETG (low-cost extrusion printers)
Part name | Qty | Notes
Cage60mm-2sides-InnerD75mm-L75mm | 2 | Laser safety snap-on cover
Mitutoyo-BD-black-ring.ipt | 5 | Only for Mitutoyo BD objectives
Adapter-MitutoyoMT1-SM2.stl | 1 | If Mututoyo MT-1 tube lens is used
3D printed parts from PA12 (SLS printer)
Part name | Qty | Notes
25mmFilterAdapter.stl | 5 | Can be also printed on hobby printer from PETG, but dimensions may be off
Large-cuvette-mount-40mm-V3.ipt | 3 | Can be machined from POM as well
sample/cuvette holders | varies | BABB/DBE resistant, autoclavable

## Options
Part # | Vendor | Description | QTY | Unit Price (EUR) | Price (CHF) | Assembly | Subassembly | Comment
Excitation laser combiner options
1008062 | FISBA | FISBA READY Beam Bio 488-520-638nm | 1 | 7400 | 7400 |  |  | Low-cost alternative, but no 561 nm line possible. Buggy software, often has troubles to start. Clunky wiring interface, need to make your own BNC breakouts.
| Cobolt | Cobolt Skyra 488-561-638, power 50-100 mW/ch | 1 | 20000 | 20000 |  |  | Alternative. Moderate-cost 4-channel laser, with mature software and good support. Requires switching mechanism or laser duplication.
| Toptica | Dual-fiber output option | 1 | 50000 | 50000 |  |  | One module serves both arms, fiber switching mechanism built-in.
Cameras
| Hamamatsu | Orca Lightning | 1 | 25000 | 25000 |  |  | SOFTWARE not yet compatible! Can be done on request. FOV 29 mm diagonal.
Detection: Olympus tube lens. Best suited for Olympus objectives, and also works well with Mitutoyo BD Apo 2X objective.
36-401 | Edmund Optics | Detection tube lens: Olympus Wide Field Tube Lens SWTLU-C f=180 mm | 1 | 525 | 525 | Detection path | Tube lens assembly | Note that Mitutoyo objectives assume f=200 mm tube lens, so magnification with this tube lens is 180/200=0.9 of nominal (eg 4.5x instead of 5x)
11-149 | Edmund Optics | Olympus Wide Field Tube Lens to T-Mount Holder | 1 | 214 | 214 | Detection path | Tube lens assembly | Olympus tube lens has unusual thread, so this adapter is handy
TMA3 | Thorlabs | Adapter with External SM2 to Internal T-Mount | 1 | 23 | 23 | Detection path | Tube lens assembly | Another Olympus tube lens thread adapter
TMA4 | Thorlabs | Adapter with External T-Mount to Internal SM2 | 1 | 23 | 23 | Detection path | Tube lens assembly | Yet another Olympus tube lens thread adapter
Lensagon1.2x-XT66P2M.ipt | Custom (3D print) | Mounting clamp (for Lensagon 1.2x, T25M-12-155I) | 1 | 2 | 2 | Detection path | Objective mount | Lensagon 1.2x telecentric lens mount. 3D printed (2 parts) and bolted together with M6. Modify the bore diameter in IPT file to fit other lenses.
Detection: Filter wheel. A servo-driven filter wheel. ONLY for telecentric lenses or long-WD objectives (> 75 mm), and 2-inch filters
5 filter positions. Motorized filterwheel is useful when multiple bandpass filters are used during an acquisition. This helps separate emission spectra of the labels.
184-902-0064-000 | Mouser | Robotis Dynamixel MX-28R | 1 | 250 | 250 | Detection path | Filter wheel servo
184-903-0163-100 | Mouser | Robotis FR07-S101K Set | 1 | 13 | 13 | Detection path | Filter wheel servo
895-USBRS485WE5000BT | Mouser | USB to RS485 Serial Converter Cable | 1 | 43 | 43 | Detection path | Filter wheel servo
FilterWheel-2inch-5pos.stl | 3D printed | 5-Position filter wheel for 2-inch filters | 1 |  |  |  |  | See folder /custom-parts/filter-wheel/
Excitation: active cooling of galvos
Active cooling is needed if galvos are run at 199 Hz and large amplitude/FOV (e.g. with 2X objective)
Fan-mount-modified-Thorlabs-CagePlate-LCP03.ipt | 3D printed | Fan mount for galvos | 2 | 5 | 10 | Excitation | Galvo assembly: OPTIONAL | 3D printed or machined
5912-412JHHU | Mouser | Fan for galvo: 40x40x25mm, 12V DC, 0.25A, dust-protected (IP rating 68) | 2 | 62 | 124 | Excitation | Galvo assembly: OPTIONAL | Cooling of the galvo is essential
varies | Hardware store | Screws for galvo fam mounting: M3, L=30 mm | 8 | 1 | 8 | Excitation | Galvo assembly: OPTIONAL | NOT available in Thorlabs M3 screw kit, need to be purchased elsewhere
ER3 | Thorlabs | Galvo Fan assembly: Rod 3-inch long | 6 | 6 | 36 | Excitation | Galvo assembly: OPTIONAL
ToDO | 3D printed | Air duct made of snap-on 60-mm cage laser cover, L=52 mm | 2 | 1 | 2 | Excitation | Galvo assembly: OPTIONAL | Same model as galvo assembly laser cover, scaled to L=52 mm before printing.
Todo | Mouser | Barrel connector for galvo fan | 1 | 2 | 2 | Excitation | Galvo assembly: OPTIONAL | Power connector for two fans
Todo | Mouser | 12V, 1A power supply, barrel jack connector | 1 | 15 | 15 | Excitation | Galvo assembly: OPTIONAL | Power supply for the galvo fans
ER4-P4 | Thorlabs | Cage Assembly Rod, 4" Long, Ø6 mm, 4 Pack | 2 | 25 | 50 | Excitation | Galvo assembly: OPTIONAL
Electronics / Signal generation
779632-01 | National Instruments | PXI-6259 (32 Analog Inputs, 48 Digital I/O, 4 Analog Outputs) | 1 | 2800 | 2800 | Signal generation |  | Only if you need more than 4 laser lines.
Alignment Tools
SI254 | Thorlabs | Shearing Interferometer with a 10-25.4 mm Beam Diameter Shear Plate | 1 | 488 | 488
SI100P | Thorlabs | Shear Plate, 5-10 mm Beam Diameter | 1 | 184 | 184
SICP | Thorlabs | 30 mm Cage System Adapter for Shearing Interferometers with Diffuser Viewing Screen | 1 | 76 | 76
CCM1-G01/M | Thorlabs | 30 mm Cage Cube-Mounted Protected Aluminum Turning Mirror, M4 Tap | 1 | 184 | 184
AC254-100-A-ML | Thorlabs | f=100 mm Ø1" Achromatic Doublet SM1-Threaded Mount ARC: 400-700 nm | 1 | 94.1802 | 94.1802
SM1V15 | Thorlabs | Ø1" Adjustable Lens Tube, 1.31" Travel Range | 1 | 36.8 | 36.8
SM1T2 | Thorlabs | SM1 (1.035"-40) Coupler, External Threads, 0.5" Long | 2 | 21.21 | 42.42
DGM05 | Thorlabs | Digital Indicator | 1 | 304.75 | 304.75 |  |  | Very useful
DIGC6 | Thorlabs | Digital Calipers | 1 | 148.21 | 148.21 |  |  | Very useful
R2L2S3P2 | Thorlabs | Grid Distortion Target, 1.5" x 1.5", 250 µm Grid Spacing | 1 | 348.45 | 348.45 |  |  | Very useful
FFM1 | Thorlabs | 30-mm-Cage-Compatible Rectangular Filter Mount | 1 | 60.95 | 60.95
DG10-1500-H1-MD | Thorlabs | Ø1" SM1-Mounted Frosted Glass Alignment Disk w/Ø1 mm Hole | 1 | 34.5 | 34.5 |  |  | Nice to have
DG20-1500-H2-MD | Thorlabs | Ø2" SM2-Mounted Frosted Glass Alignment Disk w/Ø2 mm Hole | 1 | 51.45 | 51.45 |  |  | Nice to have
CPA1 | Thorlabs | 30 mm Cage System Alignment Plate with Ø1 mm Hole | 2 | 13.8 | 27.6
LCPA1 | Thorlabs | 60 mm Cage Alignment Plate | 1 | 19.09 | 19.09
LCP01/M | Thorlabs | 60 mm Cage Plate, SM2 Threads, 0.5" Thick, M4 Tap (Two SM2RR Retaining Rings Included) | 2 | 40.51 | 81.02 |  |  | Nice to have as spares
CP20S | Thorlabs | 30 mm Cage System Iris Ø20.0 mm Maximum Aperture | 2 | 83.5981 | 167.1962 |  |  | Nice to have
SM1D12D | Thorlabs | SM1 Ring-Actuated Iris Diaphragm (Ø0.8 - Ø12 mm) | 4 | 70.104 | 280.416 |  |  | Nice to have
G061008000 | Qioptiq | Mounting Plate Without Hole | 2 | 28.75 | 57.5 |  |  | Have to be modified with SM1 threads
G061372000 | Qioptiq | Carrier FLR 40-20 | 2 | 38.8728 | 77.7456
| Fisso | Strato M-28 movable arm | 1 | 205 | 205 |  |  | Very useful
| Fisso | Thread adapter female M8/ male M6 thread reduction for Fisso | 2 | 12 | 24 |  |  | Very useful
Immesion liquids
19569/16OZ | Cargille | Immersion oil from cargille for Clarity: Code 50350 nD=1.4587 fused silica matching oil | 2 | 140 | 280 |  |  | if CLARITY is used
M3294 | TCI | Immersion oil for CUBIC: Mounting Solution (RI 1.520) [for CUBIC-R+] | 2 | 84 | 168 |  |  | if CUBIC is used
cat. no. 108014-1KG | Sigma-Aldrich | Dibenzylether | 1 | 54 | 54 |  |  | if 3DISCO/iDISCO/vDISCO used
Sample/cuvette mounting
https://github.com/mesoSPIM/mesoSPIM-hardware-documentation/tree/master/mesoSPIM_V5
https://github.com/mesoSPIM/mesoSPIM-hardware-documentation/wiki/mesoSPIM_sample_handling
Sample cuvettes | Material: glass, RI=1.52. For quartz (e.g. CLARITY imaging, RI=1.47), replace UG with UQ in the part name.
UG-205-H80 | Portmann Instruments | 10x20x80 mm | 1 | 518.5 | 518.5 |  |  | Whole-Mouse CNS Cuvettes
UG-753-H100 | Portmann Instruments | 40x40x100 | 1 | 764 | 764
UG-205-H120 | Portmann Instruments | 10x20x120 | 1 | 816 | 816 |  |  | Whole-Mouse CNS Cuvettes
UG-753-H120 | Portmann Instruments | 40x40x120 | 1 | 952 | 952
UG-751-H75 | Portmann Instruments | 20x20x75 mm | 1 | 99.5 | 99.5
UG-752 | Portmann Instruments | 30x30 | 1 | 128.5 | 128.5
UG-753 | Portmann Instruments | 40x40 | 1 | 156 | 156
UG-754 | Portmann Instruments | 50x50 | 1 | 184.5 | 184.5